## Slide 1
Optical Sensing Using fluorescence
Krishnendu Chakrabarty (PI), Richard Fair (Co-PI), Nan Jokerst (Co-PI)
PhD Students: Diksha Arora, Liji Chen, Bang-Ning Hsu, Kai Hu, Yan Luo, Matt Royal, Ozlem Senlik
Department of Electrical and Computer Engineering, Duke University, Durham, NC 27708
Research Components (Year 2)
Microfluidic Device
Cyberphysical Integration for Digital Microfluidic Biochips(NSF CNS-1135853)
Conclusions
Optical Sensor
2. Integration of IPF with qPCR Sequence Identification
3. Thin-Film Resistive Heater
Cyberphysical Error Recovery :
Physical-aware control software, on-line error recovery using sensor data
Microfluidic Device:
On-chip DNA purification and sequence identification: minimizes hands-on time
Error recovery: prevents loss of precious DNA samples due to device malfunction
Optical Sensor:
The annular thin film Si photodetectors (PDs) have been integrated with a digital microfluidic systems top plate
The PDs have a hole in the center to input optical power into the system to pump the fluorophore
Optimization of PD dimensions to achieve high SNR
The emission of light from a substance that has absorbed photons
Fluorophore of choice is 6-Hex, absorption at 532nm and emission at 556nm
PN junction Si photodetectors are used to detect fluorescence emitted by excitation of fluorophore tagged biomolecules using a laser pump excitation
Error recovery with minimum response time and no interruption of other operations
Annular thin-film PDs have been designed, fabricated, and integrated into the top plate of the digital microfluidic system. The fabricated PDs have been characterized, and have excellent dark current and responsivity, comparable to commercial thick Si PDs On-chip DNA purification and quantification have been demonstrated. Currently, the detection limit is approximately 243 copies of Alu Yb8 per 92 nl PCR mix, and the dynamic range is 103
Characterization:
Photoresponse of PD measured using 532 nm laser, coupled to multimode 200 mm core fiber
Microfluidic Device
1. DNA Purification by Immiscible Phase Filtration (IPF)
Commercially available ray tracing software, ZEMAX used to optimized PD design parameters
10 mm thick PDs integrated into a pyrex microfluidic system top plate, thin film device fabrication: p-type-SOI, diffused with phosphorous doped spin-on glass to form pn junction. Standard lithography defines dual ring ohmic contacts, mesa etched and thermocompression bonded to microfluidic top plate. Top and bottom contact leads defined
Left: Mesa etch on p-type SOI
Right: PD after mesa etch as seen under the microscope
Left: PN junction Si PD thermocompression bonded to microfluidic top plate
Right: Thin film Si PN junction PD
Test Setup
Dark Current: 97 pA
Responsivity: 0.361 A/W
4. Performance Characterization
Purification power: 1.25 x 104 with two washes
Nucleic acid retention: 32%
PCR time: 40 thermocycles in 30 ~ 50 minutes
Amplification efficiency: 76%
Sample-to-answer time: 1 hr
Cyberphysical system integration
Upper plate
Droplet
Lower plate
Computer
Electrodes
Hydrophobic Layer
Dielectric layer
Detector
Physical Subsystem
Cyber Subsystem
Algorithmic Innovations
Dictionary-based error recovery
State of the controller
State transition
Controller
Outputs: generated by the current state
S0
S1
S2
S3
S4
Input events: generated by sensing system on the biochip
Trigger rules of state transition
States
Outputs
S0
10000100...
S1
00100001...
Error dictionary
… …
Error recovery approach on cyberphysical biochip can be controlled by an finite-state machine (FSM) running in a field-programmable gate array (FPGA)
The cost and the size of the system can be reduced
… …
… …
… …
… …
… …
Dictionary item for error-free case
Dictionary item for single fault cases
Dictionary item for combination of two errors
… …
Level 0
Level 1
Level 2
A. Biochips with multiple clock frequencies
B. Operation-interdependency-aware synthesis
Synthesis method for sequencing graphs with directed tree structure
Input from the signal generator
Feedback from sensors
T phase?
Yes
No
f
f
T
f
D/M
Frequency divider
Biochemistry synthesis under completion-time uncertainties in fluidic operations
Error recovery solutions are stored as a dictionary in the memory on FPGA before the execution of the bioassay
The dictionary has a decision-tree structure
The response time of the system can be reduced to near zero
Experiment plans
Plan 1: Droplet rerouting for error recovery
Plan 2: Error recovery for mixing/dilution operations